Operating Manual
Varian SpectraAA-220

Before starting to use the instrument check the following:

1. Make sure the compressed air line is on (valve on back wall)

2. Make sure the acetylene has at least 15 psi pressure

3. Make sure the acetylene tank is turned on

4. Make sure the burner-head trap is filled with water

5. Make sure the over-flow bottle on the floor is not full

6. Be sure that the nebulizer capillary is in e-pure water 

When making solutions for any Atomic Absorption solutions, be sure to use e-pure water!

Turn on the power to the AA-220 leaving 15 minutes of warm-up time before using.

If the computer for the system is off, turn on and CANCEL at the log-on screen.

Load the software by double clicking the SpectraAA icon on the main screen

[image: image1.png]4 =181]
Fiing Dovelop | Labels | owment |
Mothods - Method 1 of 1
Worksheat Methc
TypeMode | Messuement | Opical | SIPS | Standards | GC Tess | Samplr | Motes | Cookbock |
A
b Method Select a page tab (Left) o display method parameters, or &
Tope  Flame method tab [Bottom)] to review each method. Note: Orice a
method contains resuls (indicated by a (%) in the title bar),
o Elerent [Co | Select Cetan s wil becoms dsebld. When 0C-0n (Secuence
g Wiy [ | o, o s e e
fi @ 1e Units ma/L.
] 4 Sampling Mode- ‘Elame type & Gas flows (L/min]
Y- Flame Type |
] [ N20 Flow =]
5 € Miero Samping Agalylene Flow 6.35
4 o =
Instument Mods I UsesiFs
@ Absorbance
€ Enision
< Back Hext>
@
Ok Cancel Help.
Goude
Astar|| @ & A ¥ ||[S5poctan 71 AddMethod - Paint B a0 5 1049M



The software should open to a screen like below:
Everything within SpectrAA is done in Workbooks.  Double click the Workbook icon.
The following screen should appear:

[image: image14.png]SpectiAA 220 =18 %]

Help

1A

ANARN

Worksheet Oriented AA Award Winning

Spectroscopy Software

| Worksheet Bepots ‘ Admirisation
[Password Protecton: OFF]

A 4

st | & G 4 ¥ [ specran == 8 223PM

(c) Copyright VARIAN AUSTRALIA 1997




Fill in the information for your new worksheet.  You need to enter a Name and the number of samples.  It is suggested that you enter more samples than you expect to run because you can not add more samples once you have started your analysis.  Click OK when ready to move ahead.

The following screen should appear:

[image: image15.png]SpectiAA 220
Help

=181]

Spefia A B3

New Worksheet
Watksheet Detais

R T
Soes [0 =

Yo must enter a name and the numbe of samples when
creating a niew worksheet. Al ther parametes ate optional The
actual DOS flename is created automatialy by the system

[ Cancel Help

=3 i

| Worksheet Bepots Admirisation

il

Esit

[Password Protecton: OFF]

A 4

Award Winning
Spectroscopy Software

(c) Copyright VARIAN AUSTRALIA 1997

T e

BT 2




You then want to add a method(s) for your analysis.  Click the Add Method icon and the screen below should appear.
[image: image16.png]SpectiAA 220 - Calcium Analy:
Options  Exit  Help.
e Devop | Lobes | nswnent |

Warksheet Methods T

Add
Methods

Esst
Sequential.

Edt
Methods.

Edt
Sequence
Parameters.

Help

SIPS offine

Astat|| & G A 9 [[[S spection S Nemeworksheet - Pain BEi s 2um




You will want to load from the Cookbook initially, until you have created a Method Library.  We can use only flame, and you can scroll through the displayed methods to find the metal ion you want to set up the method for.

Double click the metal ion and the following screen should appear:

[image: image17.png]4 =181]
Firg Devep | Labok e
Worksheat Methods e T
- . ] Review
O - i i cthod: 7]
Selec one o ore methods o et he cockbock o the metd ey to dd o
yous cuent st of watksheet methacs,
Fast
Seauential Losd From Search
' Cookbook Element
Edit € Method Liprar, atin
@ 2. Method Libary M
= Elament Matis
Method Type
7 Flame
Edit - WV E
Sequence zx [~ Fumace
Parametrs. = Ve
zx [ Zeeman
Cancel Help.
Goude
st | @ &4 ¥ [ specran §IDevelapSereen - Paint 1B a0 8 1048M




In this screen, you develop the method for the metal of choice.  In the example Ca will be analyzed.  At this point you will need to edit the generic method for the analysis.  Double click the Edit Method icon.  The method set up screen will open
[image: image19.png]SpectiAA 220 - Chem336 ME]
Instument Edt Yiew | Opons st Ho
Fing | v IrackAnaysis  CulT L nstmert |

S| 264 | 0.0003
— >

v Search
| et ,

1 Impart Marks.
Export Matks.

abs Sample 001
o0

Select Piter.

Sample 005 01800
Somple 005 01788
Somple 007 00534
Somple 008 00540
Somple 003 00020
Sample 010 00008 100 150

bo E
Sonple 01T |i e Oty Tne [ Edwsls

10:39:52 AN 05/10/2005
11: samle ol 10:39:52 M 05/10/2005

0005

0000
0002

1
1

Fa (Flame) - analysis cime to dave: Sminuces
jorksheet Saved

bnalyse pressed scop ~
il |

Method 1 of 1 =Ca Lamp 1
Absarbance. Scale Expansion Manual Fow 1 of 50
Inegiale 505 Aidcelylene BCOif Resul ows: 10

Piintthe contents of the data log

st | @ &4 ¥ [ specran §1ViewMenu - Paint B a0 5 10:46M




The Tabs that you will modify are the Type/Mode, Measurement, Optical and Standards.
In the Type/Mode tab, set the Units to the units you are going to use (mg.L, ppm, Abs), sampling to Manual and the Flame type to Air/Acetylene like in the figure below then click Next which will open the Measurement tab
[image: image2.png]4 =181]
Fiing Dovelop | Labels | owment |
Mothods - Method 1 of 1
Worksheat Methc
TypeMode | Messuement | Opical | SIPS | Standards | GC Tess | Samplr | Motes | Cookbock |
A
b Method Select a page tab (Left) o display method parameters, or &
Tope  Flame method tab [Bottom)] to review each method. Note: Orice a
method contains resuls (indicated by a (%) in the title bar),
o Elerent [Co | Select Cetan s wil becoms dsebld. When 0C-0n (Secuence
g Wiy [ | o, o s e e
fi @ 1e Units |abs
] 4 Sampling Mode- Elame type & Gas flows (L/min]
Py Flame Type |
] € pusoroml AiFlow ==
H  Mito Samping - [ew0 =
Instument Mods I UsesiFs
@ Absorbance
€ Enision
< Back Hext>
@
Ok Cancel Help.
Goude
st | @ &4 ¥ [ specran SFEdiMethod - Paint 1B a0 5 1050aM




In the Measurement Tab, the only mode available to the SpectrAA-220 for Measurement Mode is Integration mode.  Under Calibration Mode a number of options are available depending upon the analysis you are doing.

· If you want to simply measure absorbance and do the data analysis on your own, then select Scale Expansion
· If you have a set of standards of known concentration and want the instrument to do the calibration curve – giving results for your samples in the units of concentration select the Concentration button.
· If you are doing a standard addition experiment, selecting Standard Addition, provides you the chance to generate the standard addition curve automatically.
Depending on the Calibration Mode used, you will then have a chance to select numbers of replicates for both sample and standards.  The Time(s) simply provides a means of changing the integration parameters used for measuring the data.  

[image: image3.png]@ =)

e Devop | Lobes | nswnent |

Methods - Method 1 of 1

WakshestHethc
Typeiode  easierent | Opical | SIPS | Standtds | GCTests | Sanpler | Notes | Cookbook |

a
o Messurement Mode Calbration Mode:
€ PROMT € Concentation
£ & Inegraion € Standard Addiions
€ Peak Height € Brackeling Standards
1. € Peakdvea @ el il
L | Replicates
Miinum Resgng  [Foo00=] | Sterdd =|
i Sample =
gy | smoonng Spant
|5 e
Measuemert 50 = =
ReadDelay = =

<Back | New>
- Ca

ok Corcl |tk

GCUide

st | @ &4 ¥ [ specran SFEdiTypeMode - Paint B a0 5 1051 AM




Scale Expansion Mode – Used when measurements of absorbance are desired.

[image: image4.png]@ =181
Fing Dovelop | Labels | owment |
Watksheet Methe
TopeMode  Messuement | Opical | SIPS | Standards | G Tess | Samplr | Notes | Cookbock |
A
M Measurement Mode Calibraion Mode:
« PROMT & Concentation
g & fnitearafior; € Standard Additions
C Pesk Height  Brackeling Stancrds
1. € Peskirea € Soale Expansion
FE ) R e
L | Replicates
Minimum Reading [0.0000=] St =
d Sample =
&[5 s [Gpom =1
G Time (5)
Messuement  [50 = =
ReadDelay = =
<Back Hext>
R
ok Corcl |tk
st | @ &4 ¥ [ specran B8 1124 AM




Concentration Mode – Used when you have a set of standards and want the instrument to report results directly in concentration units.  You will want to go back to the Type/Mode tab and define the units you are using (gm/mL, Abs, ppm).
[image: image5.png]4 =181]
= Deveon | Labels | instument |
Watksheet Methe
TypeMode Measutemert | Optcal | SIFS | Stondads | OC Tests | Samplr | Notes | Cookbook |
A
M Measurement Mode Calibraion Mode:
 PROMT € Concentration
g & fnitearafior; ' Standard Additions
€ Peak Height € Biocketing Standards
1. € Peakdvea € Seal Expension
8 A Repicates
Minimum Reading [0.0000=] Sl =
¢ Sample = |
ﬁﬂj-"\ £ Smoohing 5pont <)
i Tine (s)
Measuemert 50 = =
ReadDelay = =
< Back Hext>
~Cs
ok Corcl |tk
SIPS affine
Astat|| & £ A ¥ ||[S5pectan f]Methodédded - Paint BEGDD 23ru




Standard Addition Mode – Used when using the technique of standard addition.
The Optical Tab allows you to set up the lamp that is to be used.  Some lamps (Varian) will be identified by the instrument and set themselves automatically.  Other lamps will need you to set the position, current (suggested on lamp), wavelength and slit width.  The default settings are usually a good place to start. 

[image: image6.png]@ =)

= Deveon | Labels | instument |

WakshestHethc
TypeiMode | Measusment Optical | 51P5. | Standacs | 0C Tests | Samplr | Notes | Cookbook |

A
W
E HE Lamp
5 I8 i
= Lomp Posiion
E
M Lamp Curent (mé)

Manachiomtor

o | Wavdlergh o) @ =
o | Skwidh ) o5 o
Backguound Corecton [BC1 5

ok Corcl |tk

GCUide

st | @ &4 ¥ [ specran S EdiMeasurement - Paint B a0 5 1052aM




If using the Concentration or Standard Addition methods, then the Standards Tabs will allow you to enter the concentrations of the standards into the instrument so that once measured, the curves can be generated.

[image: image7.png]- =18l %]
Fiing Deves | Lsbek | lnsmmen |
Wolksheet Methc
TypeMods | Masssement | Dptca | SP5 St | 0 Tews | Sampe | Notes | Cookbeok |
s Smad  Coc  Refabs
M [STANDARDT 02500 InstZeo e b =
STanaRD 2 05000
o || B =
A e Rscaibraion Rate o0 =
=S STANDARD 5 Reslope Rate 5o =
& |5 oo kst =
= - [STANDARD 7 Expansion Factor 1.0
sTanDaRD 8
=
) S TANDARD 9 Conc Decimal Places =
liﬂ-"\ 5 [stanoarD 10 Calt, Hlgoifm Newhaiond
P sipelens Lower & Upper Link
Cobaton [0 = %[00 =% =
Reslope 750 =% [1250=] % =
< Back Hext>
e
ok Corcl |tk
st | @ &4 ¥ [ specran §f15tandardsMeasurement 1B a0 8 11:25M




Concentration Mode
[image: image8.png]@ =)

e Devop | Lobes | nswnent |

Methods - Method 1 of 1

Warksheet Meth:

Typeiode | Messement | Opicl | SIPS  Standatds | € Tets | Samper | Notes | Cookbeck |

Standards __Cone Ref Abs

=

InstZero Rate 0

ik

i

El

E

Expansion Factor [i

Conc Decima Places

Slope Tests Lower & Upper Lirit

Calbration 200 = % [1500=] %
Resiope 750 =% [1250] %

<Back | New>
- Ca

al

ok Corcl |tk

SIPS offine

st | & 5 A Y [|[Sspecton 7]Standardéddiion - Paint PGS 23




Standard Addition

Once you have set all of these parameters, click Ok and the parameters will be set in the instrument.

Click on the Instrument Tab and then the Optimize button and the following screen should open.  
[image: image9.png]4 =181]
[ wam [ T swnent |
Exp s RSD Mesn Abs
Seleet ‘ Flame Optimization
T 5o Optimization: Lamp. Gias Flow [L/min imple 001
Tube La HC Lamp Ovidant Acetylene
1 |SampleC 1307 20.00] 8.00]
e Oplinize Lamps
S IREEEY | 65 Optimize Signal 18.00]
4 [sampet s
EMETRE Rescale 16.00]
- 6 [sampet
7 [sampleC | 050 It Zen0 400
8 [Sample [ AT
9 [sampet
10_[SampleC | 000, Gan 2 % 11.00] 1.50) 180 150
T [samet | g gog e oo = Tne [ AReE
Hoo | metver press |
pualy=is Tk Buner Heig 00 =] mm
tethod signal B B =
Sersiviy Check [0 ma/L gves about 02 Abs o 422.7 m, N20/A bumer
bopesnize stop
R ok Carcel Help =
i ]
Method 101 == Lanp 1
Asoance Sedle Expansion Hanal Fowd of 50
iteqae 505 Ailhcabiene BCOf Rt rows: 0
Astar|| @ & A ¥ ||[S5poctan SPEdiCookbock - Paint 1B a0 5 1053AM




This screen gives you the chance to optimize the instrument before doing your analysis.  The parameters that need to be optimized are:

· Lamp position – click on Optimize Lamp and maximize the signal displayed

· Flame position – click on Optimize Signal and maximize the absorbance measured

· Gas Flow – changing ratio of Oxidant to Acetylene to maximize the absorbance measured

Lamp Position 

The lamp position is optimized by selecting the Optimize Lamp button and adjusting the two small nuts on the bottom of the lamp.  You want to maximize the emission signal that is reaching the detector.  This should be done without the flame being ignited,
Flame Position

To optimize the flame position, select Optimize Signal, ignite the flame and aspirate a solution with your analyte into it (usually the most concentrated standard) and adjust the burner-head position (up and down, in and out) using the two knobs at the base of the burner.   Again you are trying to get a maximum absorbance.

[image: image10.png]4 =181]
[ wam [ T swnent |
Exp s RSD MeanAbs
Seleet ‘ Flame Optimization
T Sa |- Optinizaton Sinal GasFlow L/} e 001
iz La | FG Signal Didant Acellens
1 [SempleC | 130 20] e
e Oplinize Lamps
S IREEEY | 65 Optimize Signal 18.00]
4 [SampleC so0]
5 [SempleC .
- 6 [SempleC
7 [sampleC | 050 It Zen0 400
8 [Sample [ AT
9 [SampleC
10_[SampleC | 000, Gan 21 % 11.00] 1.50) 180 150
T fsamet | g o0 e oo = Tne [ AReE
HoP | brarver press =
pualy=is Tk Buner Heig 00 =] mm
tethod signal B B
Sersiviy Check [0 ma/L gves about 02 Abs o 422.7 m, N20/A bumer
bopesnize stop
R ok Carcel Help =
i ]
Method 1 of 1 <Ca Lanp 1
Absobence Sedle Expansion Hanal Row of50
iteqae 505 Ailhcabiene BCOf Resul ows. 0
Flame actve
Astar|| @ & A ¥ ||[S5poctan FlnstumentDptinizeLamp B a0 5 1056 M




Gas Flow

Usually the default settings set by the Cookbook give good results.  Varying the ratio of Oxidant to Fuel can improve your Signal slightly.

This optimization of the signal is an iterative process where you will optimize position, the fuel ratio, and then re-optimize position, etc.  
Once the system has been optimized, click Ok and get back to the Instrument Tab.  You are now ready to measure your samples.  Press Start and follow directions on the screen.  Initially you will be aspirating e-pure water.  You will want to aspirate e-pure while zeroing the instrument. (NOTE:  You should aspirate solution all the time the flame is on.)
Aspirate your sample into the flame and press the Read button, and continue with all of your samples.  You should generate a view like the following.

[image: image11.png]SpectiAA 220 - Chem336 L=18]
Instument Edit View Options Ext Help

Fing | Devekp | Lsbels [ smmer ]
o4 || -0.1715 00 | -0.1715
u Sample. Abs, ‘Sample 004
§ 2 [Sample 002 01714 200
e 4 |Sample 004
D | C o — 100
7 [seonceo (N
! 8 [soncie0vs (N
o i . | 0
Yo | [ oo p— Ty £ FE PR,
T [sampeorr [ | ey —— Tine [ A
) 1: Sauple 001 -0.1700 0.4 -0.1700 -0.1693 -0.1633 -0.1707 N
10:57:46 AM 05/10/2005 o
| |
Method 1 of 1 —~tCa Lamp 1 Autorun in progress.

st | @ &4 ¥ [ specran Funtiled - Paint 1B a0 5 1059AM




If using the Concentration or Standard Addition Mode, you will sometimes find the output to look like:

[image: image12.png]SpectiAA 220 - Calcium Analy:
Instument Edt View Dptions  Ext

Fing | vewks |

Help

Labes e

Concentraion %RSD

Mean Abs

| UNCAL 36.9

Select

-0.0004

Sample
e Labels

Sample 001
Somple 002
Sample 003
Somple 004
Sample 005
Somple 005
Somple 007
Somple 008
Somple 003
Somple 010
Sample 011

P P Y O =P

H

ks
260

Sample 002

200

1.00

000

bo
Zoom |

F)
ovety |

100
Tine:

150

Autoscale

[rastrument zero (Postreat

0001abs, Gain=22¢]

rve  sawple
1 Sample 001

Cone na/L
AL

sxsp
6.5

4502: Sample uncalibraced
2 sauple 002

HMean abs
-0.0004

Readings
-0.0006 -0.0004 -0.0003
2:38:02 P 05/10/2005

2:38:02 m 05/10/2005

Methad 1 of 1
Absarbance.
Inegiale 505

=Ca
Standerd Addiions
Ailhcelylene.

Lanp 1
Manual
BCONF

Autorun in progress.
Sample 002

Warksheet Saved

Astat|| & G A 9 [[[S spection )StendardsinStandadd - Pai

Preread delay 35

238PM





The UNCAL indicates that the sample was outside of the calibration standards.

When you are finished, the output should look something like the following:

[image: image13.png]SpectiAA 220 - Chem336 L=18]
Instument Edit View Options Ext Help

Fing | Devebp | Labels et |
Epabs s Mean abs
skt ‘ 0.0003 26.4 0.0003
VT Sawe s Sange 01
Tube | Lobek Hiet
T [sempkot o000 :
7 [senpe00z_[n.0001 I
3 [senpk0m 00002 i
4 [sonpk0ot 0178 1 o
5 [sonpk0s 01600 I
6 [senpkse 01788 f
7 [senpkoor 005 ;
o [senpk0s8[00540 |-
9 [senpk08 00020 | Som
10 [Sanpe 01000008 I T % o =0
T [sanpe 01T =i 75 oy Tne [ Edwsls
Toro9:52 & 05/10/2005
11: Sample 011 10:39:32 AM 05/10/2005 o
o (Fianel - anaiysis ine to dave: Swimures
oxtshect Saved
finalyst pressed stop =
KIN| |
Mebod 1ol s Larp 1
Absoibance SealeEvpansion e Fow of50
Inegae50s  Auvhcobions B0 Rosa ows. 10
Goue

Astar|| @ & A ¥ ||[S5pectian 7]untted - Paint 34

ERLTTY




To print out the results, go under Options on the Menu  bar[image: image18.png]Dptions Ext_Help

Fiing Develop

Warksheet Methods

T |

Element Matis

Revien
add (S
Hethods.
U

East Down

Sequental (S
= Delete

Edil L= |

Methods.
= Copy toLibray.

£l

Sequence

Parameters

el ide

Asn| | @ €AY

@ Srectiaa Pt Pt

B 1047 aM




, then to Datalog, to get to the print function.

A table which contains all of the data you collected will be printed.  The table will contain the measured values for each replicate, an average as well as the relative error.  It will also give the time and date for the analysis and the metal being analyzed.  You can also export the document as a text file, though direct import into Excel is not simple.
PAGE  
13

